| Name | Security Level |
| --- | --- |
| BELLSING® JOYLIVE Series BMBA Receiver | Open |
| Version | A Total of 7 Pages |
| REV X1 |  |

BVA350X24002

Product Data Sheet

| Prepared by: | Quanlong.Lv |  | Date: | 2022-12-16 |
| --- | --- | --- | --- | --- |
| Auditor: | Nick Zheng |  | Date: | 2022-12-16 |
| Approver: |  |  | Date: |  |

Shenzhen Bellsing Acoustic Tech. Co., Ltd.

Product Specification

Product number :  See below List

| Model Number | Description |
| --- | --- |
| BVA350L24002 | SIZE 2, MEDIUM, LEFT, NYLON TUBE |
| BVA350R24002 | SIZE 2, MEDIUM, RIGHT, NYLON TUBE |

1. Description and Application

1.1 Description

RIC Module for hearing aid units

1.2 Application

Medical equipment

2. Mechanical Layout and Dimensions

2.1 Dimension

2.1.1 Mian Dimension

2.1.2 Size Dimension

BVA350L24002, BVA350R24002

2.2 Material List

| QTY |  | DESCRIPTION |  | ITEM |  |
| --- | --- | --- | --- | --- | --- |
| 1 |  | Nozzle |  | 1 |  |
| 1 |  | BRC350 Receiver |  | 2 |  |
| 1 |  | Top case |  | 3 |  |
| 1 |  | Bottom case |  | 4 |  |
| 1 |  | Nylon11 Tube, OD 1.0mm |  | 5 |  |
| 1 |  | Connector rear housing |  | 6 |  |
| 2 |  | Metal Sleeve |  | 7 |  |
| 1 |  | Connector |  | 8 |  |
| THIS BILL OF MATERIAL IS PRIVIDED AS A REFERENCE, BELLSING CONTROLLED BOM TAKES PRECEDENCE. |  |  |  |  |  |

3. Electrical and Acoustical Specifications

3.1 Frequency response

Typical frequency response measured according to chapter 3.3 and 3.5

An undamped magnetic balanced armature receiver with an extended high frequency response.

| Test Item |  | Nominal | Tolerance |
| --- | --- | --- | --- |
| Sensitivity(dB) | @200Hz | 109.0 | +/- 3.0 |
|  | @500Hz | 107.0 | +/- 3.0 |
|  | @1000Hz | 106.0 | +/- 3.0 |
| Peak1 | Frequency (Hz) | 3000 | +/- 350 |
|  | SPL(dB) | 114.0 | +/- 3.0 |
| Valley | Frequency (Hz) | 5600 | +/- 500 |
|  | SPL(dB) | 97.5 min |  |

3.2 Total Harmonic Distortion (THD)

Typical THD measured according to chapter 3.3 and 3.5

Total Harmonic Distortion

Device Will Not Exceed Total Harmonic Distortion Levels Listed Below.

| Test Item |  | Nominal | Comments |
| --- | --- | --- | --- |
| THD(%) | @1/2 1st peak | <5.0 | AC drive 0.16 Vrms |
|  | @1/3 1st peak | <5.0 |  |

Electrical

| DC Resistance @20℃ |  | 48.0Ω+/- 10% |  |
| --- | --- | --- | --- |
| Impedance @500Hz |  | 72.0Ω+/- 15% |  |
| Impedance @1000Hz |  | 115.0Ω+/- 15% |  |

3.3 Test Conditions

| Nominal Source Voltage |  | 0.16 Vrms，0 mA DC BIAS |  |
| --- | --- | --- | --- |
| Source Impedance |  | ＜1 Ohm |  |
| Tubing |  | NONE |  |
| Coupler Cavity |  | 1.26cm³，Simulated IEC711 Coupler（IEC 60318-4） |  |

Unless otherwise specified the range of atmospheric condition to complete Measurements and tests will follow 23±5℃ ambient temperature, 25~65% relative humidity

Operating Temperture: Sensitivity at 500Hz will not vary more than +1/-3dB from -17℃~63℃

Storage temperture: -40℃~63℃

Solder Type: SAC305

3.4 Measurement setup

Measurement signal:

Logarithmic sine sweep, 1.5s, 20Hz~20 kHz

No RF chokes included in the measurement!

All acoustic and electric measurements at 23±2°C

3.5 Measurement adapter