| Name | Security Level |
| --- | --- |
| BELLSING® Joylive Series AMBA Receiver | Open |
| Version | A Total of 4 Pages |
| REV E |  |

BRC350H32667

Product Data Sheet

| Prepared by: | Shane Wu |  | Date: | 2021-07-28 |
| --- | --- | --- | --- | --- |
| Auditor: |  |  | Date: |  |
| Approver: | Frank Huang |  | Date: | 2021-07-28 |

Shenzhen Bellsing Acoustic Tech. Co., Ltd.

BRC350H32667 AMBA receiver

# Description

## Advanced Miniature balanced-armature（AMBA） receiver with magnetic-radiation shielding for use in ITE (In The Ear) , CIC (Complete In the Canal) and IIC (Invisible In the Canal) & In-ear Monitor applications with standard response. This transducer can also play an excellent role in commercial hearing products,TWS applications.

| ## Name: | ## Joylive Series AMBA Receiver |
| --- | --- |
| ## Part Number: | ## BRC350H32667 REV E |

# Construction Specification

## Weight:  0.16 g

## Mechanical dimension：please refer to the drawing below

*A positive voltage applied to the negative terminal () will result in an increase in pressure at the sound outlet.

# Features & Benefits

## Perfect package size for ITE, CIC, IIC & headset, TWS applications

## Outstanding efficiency

## Substantial maximum output

## Magnetic-radiation shielded

# Electro-acoustic Specifications

## Test condition

## Acoustic tubing: 10.0mm of 1.0mm diameter tubing

## Acoutic coupler: IEC711 coupler

## Constant voltage drive: 0.287V RMS

## Environment: 23℃, RH 50%

## Acoustic Parameter I

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
| Sensitivity(dB) | @200Hz | 110.0 | +/- 3.0 |  |
|  | @500Hz | 107.5 | +/- 3.0 |  |
|  | @1000Hz | 106.0 | +/- 3.0 |  |
| Peak1 | Frequency (Hz) | 2800 | +/- 350 |  |
|  | SPL(dB) | 116.0 | +/- 3.0 |  |
| Peak2 | Frequency (Hz) | 5600 | +/- 500 |  |
|  | SPL(dB) | 111.0 | +/- 3.0 |  |
| Maximum output@Peak frequency (dB) |  | 124.0 | N/A | THD<10% |

## Acoustic Parameter II

| Test Item |  | Nominal | Comments |
| --- | --- | --- | --- |
| THD(%) | @1/2 1st peak | <5.0 |  |
|  | @1/3 1st peak | <5.0 |  |
|  | @500Hz | <5.0 |  |

## Electric Parameter

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
| Impedance(Ω) | @500Hz | 220.0 | +/- 15% |  |
|  | @1000Hz | 350.0 | +/- 15% |  |
| DCR(Ω) |  | 156.0 | +/- 10% |  |

## Nominal Frequency Response Curve

# Package and Storage Condition

## Package: PET bubble pack, 1 pcs per bubble, 100 pcs per pack

## Storage temperature range: -40℃ ~ 60℃

## Storage humidity range: <60%